XRDTIME

A program for the collection of X-ray diffraction data
with variable step counting times

I.C. Madsen and R.J. Hill
CSIRO Division of Minerals
P.O. BOX 124 Port Melbourne   3207
Victoria 
Australia

Phone   +61 3 9647 0366
FAX     +61 3 9646 3223
Email   Ian.Madsen@minerals.csiro.au

## Introduction

This document describes the background to, and the operation of, a method of data collection for X-ray diffraction data using variable step counting times. The program is distributed freely, but the work of the original authors should be referenced with any published results.

## The Problem

For X-ray diffraction data collection, a number of factors act to progressively decrease the observed intensities as the Bragg angle increases.

The major factors causing this decrease are:-

- Lorentz-polarization factor

- Thermal vibration

- Atomic scattering factor

However the fall-off is partially compensated for by the effect of:-

- Reflection multiplicity

- Sample absorption (for cylindrical geometry)

The net result, however, is a fall in diffracted intensities that can be larger than two orders of magnitude across the pattern.

In spite of this intensity decline, diffraction data for Rietveld analysis is universally collected with the same counting time (or monitor count) for each step in the pattern. As a result, the peaks at high angles are collected with much lower precision than those at low angles. This is unfortunate since the mid to high angle regions of the pattern contain more information (peaks) than the low angle regions.

## The Cause

The fall-off in the intensity of a diffraction pattern can be represented as follows

I   µ  ƒ2  x  Lp  x  TC  x  m  x  A  	 (1)

where	ƒ2 	is the scattering factor
	Lp	is the Lorentz polarization factor
	TC	is the thermal correction
	m	is the reflection multiplicity
  and	A	is the sample absorption factor

### Lorentz-polarization factor

The Lp factor is the dominant effect in terms of variation as a function of 2q, exhibiting a range of about 400 across the pattern.  The additional polarization correction introduced if a crystal monochromator is used only slightly reduces this variation.

### Thermal correction

The thermal motion of atoms in the structure also serves to decrease the intensity of the high angle regions of the pattern, although the magnitude of the effect is less than the Lp factor.  For B = 0.5Å2 ( a value found in typical oxide materials) the variation across the pattern is about 1.5.  For larger values of B (3.0Å2), the correction factor can vary by as much as 10 between the two ends of the pattern.

### Scattering factor

While the overall magnitude of the scattering factor is highly dependant on the average atomic number, the variation between the low and high ends of the pattern is much less dependant.  A decrease in intensity of about 3 to 5 times is typical for a wide range of compositions. Note that XRDTIME does not take the absolute magnitude of scattering into account - only its variation as a function of diffraction angle.

### Reflection multiplicity

The decrease in pattern intensity as a function of 2q exhibited by the preceding factors is partially compensated for by the effect of reflection multiplicity.  On average, the reflections at high angles will have higher multiplicity than those at low angles.  In addition this difference will be largest for the highest symmetry.  An intensity increase of 6 to 8 times for cubic and 1 to 2 times for monoclinic materials can be expected.  The correction applied has the form

m = a + b ln2q 	  (2)

where m is the average reflection multiplicity and a and b are variables. See Madsen and Hill (1994) for details of the method of determination of these values.

### Absorption factor

In conventional Bragg-Brentano diffraction, sample absorption decreases the intensities of all peaks equally, that is, there is no variation in peak intensity as a function of 2q. However, if cylindrical (capillary) geometry is used (Hill and Madsen, 1991), the affect of absorption of the beam in the sample is to increase the diffracted intensities at high 2q relative to low 2q.

The correction is dependant on the linear absorption coefficient (m) and the radius of the tube (R). For example, for rutile (TiO2 ) packed in a 0.5mm diameter tube, and using CuK radiation, mR is about 4.0.  These values lead to an increase in the high angle intensities by a factor of 12 relative to the low angle peaks.

## The solution

To compensate for the overall decrease in pattern intensity, an algorithm has been devised to calculate the optimum step-counting time, taking the afore-mentioned parameters into account .  Equation (1) is used to calculate the expected intensity at each point in the pattern and the counting time is adjusted accordingly.  The strategy uses shorter counting times at low 2q, where diffracted intensities are high, and longer times at high 2q to compensate for the lower intensities.  Further details of this variable counting time (VCT) algorithm can be found in Madsen and Hill (1992) and Madsen and Hill (1994).

## The Program

XRDTIME is a program, written in FORTRAN, for the calculation of variable counting time data collection conditions. The program is ‘default driven’, that is, the most commonly used conditions are suggested and will be accepted by simply pressing the enter key. The details of the program input are given below:-

| Program Prompt | Meaning |
| --- | --- |
| Note - the value displayed after the [D: will be accepted if no entry is made. |  |
| Enter START angle (deg) [D: 10 > | The lowest 2theta angle to be collected |
| Enter STOP angle (deg) [D: 160 > | The highest 2theta angle to be collected |
| Enter STEP WIDTH (deg) [D: 0.020 > | The increment between data points in 2theta |
| Enter BLOCK WIDTH (deg) [D: 1 > | The width of the data blocks. The counting time will be changed after this many degrees. |
| Enter TOTAL COUNTING TIME (sec) [D: 14400 > | The total data collection time for the data set in seconds (do not enter the number of seconds per step) |
| Enter MINIMUM step count time [D: 0.20 > | By default, the program assigns 0.1 of the total counting time as a fixed time across the entire pattern. This ensures that adequate statistics are accumulated in the background regions of the low angle regions. The remaining available time (0.9 of the total) is allowed to vary in accordance with the VCT regime. |
| Enter TUBE TARGET Fe Co Cu Mo Ag [D: Cu > | Used to define the wavelength for this data collection |
| Is a monochromator fitted Y/N [D: Y > | Used to modify the Lp correction for the presence of a crystal monochromator in the diffractometer system. The use of a curved graphite, diffracted beam monochromator is assumed. |
| Enter THERMAL VIBRATION (B) value [D: 0.5 > | The average thermal parameter for the material being examined. Typically 0.3 - 0.5 for oxide materials. Values of 2.0 - 3.0 might be acceptable for organic materials. |
| Enter uR for DEBYE-SCHERRER geometry 0.0 for flat plate [D: 0.0 > | The product of the linear absorption coefficient, the radius of the sample and the packing density if capillary geometry is used. typical values for packing density are 0.3 - 0.4. Enter 0.0 for flat plate geometry. |
| Enter code for CRYSTAL SYSTEM 1 = Cubic 2 = Tetragonal 3 = Orthorhombic 4 = Hexagonal 5 = Monoclinic 6 = Triclinic [D:3 > | A code for the crystal system of the dominant phase in the sample. |
| Enter CHEMICAL FORMULA ? Y/N [D: N > | The chemical formula is used to calculate the scattering factor curves for the sample. An ‘N’ answer will assume a composition of Fe (an intermediate value of scattering for many materials). See the next box of the answer is ‘Y’ |
| Enter SYMBOL for element # 1 (Return to end list) > Mg Enter SYMBOL for element # 2 (Return to end list) > Al Enter SYMBOL for element # 3 (Return to end list) > O Enter SYMBOL for element # 4 (Return to end list) > Enter number of MG atoms in formula [D:1> 1 Enter number of AL atoms in formula [D:1> 2 Enter number of O atoms in formula [D:1> 4 Atom type MG AL O # of atoms 1.00 2.00 4.00 | Enter the symbol for the each of the elements in the compound (15 maximum). Terminate the list by pressing enter with no entry. Then enter the number of each of the atoms in the formula. The entered composition will be displayed and used to calculate the average scattering factor curves. |
| Do you wish to write these conditions to PW1710 compatible files Y/N [D:n> | Enter ‘Y’ if you wish to write a file of commands to control the diffractometer. Note that the commands currently support the Philips PW1710 and PW3710 controllers. Modify subroutine WRITE to support your own diffractometer controller. The file is called COMMAND.LIS |
| Do you want to collect FIXED TIME data Y/N [D:Y> | You can write the diffractometer commands to collect a fixed count time (FCT) data set after the variable count time (VCT) data set by entering ‘Y’. The FCT data will be collected with the same total count time as the VCT data. |

## Files supplied

The files supplied with XRDTIME are:-

XRDTIME.FOR	the source code for the main routine for the program

BLKABS.FOR	the source code for the capillary absorption correction routine

SCATFAC.FOR	the source code for the scattering factor routine

MAKEFILE	a makefile for compilation and linking of the three routines (setup for WATCOM FORTRAN77, but could be translated to other compilers.

XRDTIME.EXE       the executable version of XRDTIME for PC

XRDTIME.OUT	a typical output list file

COMMAND.LIS	a typical list of Philips PW1710/PW3710 diffractometer controller commands

DOS4GW.EXE	a DOS extender required to run XRDTIME on the PC. Place this file anywhere on the path.

REFRMVAR.FOR	a program for reformatting the raw data returned by the (PW1710/3710) diffractometer controller. This will have to be modified if other controllers are used.

VCTDISTR.DOC	this document in Word for Windows 6.0 format

## Typical output file

The following is a typical output file (XRDTIME.OUT) from the XRDTIME program

XRDTIME - calculation of variable count time data collection conditions

Calculation run at  16:04  on  15/03/1996

All published use should refer to

# Madsen and Hill, (1994), J.Appl.Cryst., 27, 385-392 #

Start = 10.000    Step =  0.020    Stop = 160.000     Block width (deg)=  5.

Total count time = 100000.0 sec      Minimum step count time =  1.35 sec

Step count time for fixed time data collection = 13.35

CU target tube - wavelength = 1.54056

Monochromator correction INCLUDED in LP

Average temperature factor of  0.50 assumed.

mu-R used for absorption correction = 0.000

Crystal system used for multiplicity effect is Cubic

Formula used to calculate scattering factor curves =

Atom type    MG    AL    O

# of atoms  1.00  2.00  4.00

Block  Angle - Limits    Avg.     Avg.     Avg.   Avg.    Avg.    Average

#    Low      High      LP      Scatt    Temp   Trans   Mult.   Int./1000

1   10.000   14.980    78.08  4287.53   0.995   1.000  14.663     4884.23

2   15.000   19.980    38.62  3855.52   0.990   1.000  17.389     2563.81

3   20.000   24.980    22.85  3421.22   0.984   1.000  19.415     1493.85

4   25.000   29.980    14.99  3022.05   0.976   1.000  21.030      930.14

5   30.000   34.980    10.51  2669.72   0.968   1.000  22.373      607.34

6   35.000   39.980     7.73  2362.62   0.957   1.000  23.522      411.06

7   40.000   44.980     5.89  2095.23   0.946   1.000  24.527      286.19

8   45.000   49.980     4.61  1862.21   0.934   1.000  25.420      204.00

9   50.000   54.980     3.71  1659.24   0.921   1.000  26.223      148.50

10   55.000   59.980     3.04  1482.70   0.907   1.000  26.953      110.30

11   60.000   64.980     2.55  1329.38   0.893   1.000  27.622       83.59

12   65.000   69.980     2.18  1196.39   0.878   1.000  28.239       64.70

13   70.000   74.980     1.90  1081.09   0.863   1.000  28.813       51.21

14   75.000   79.980     1.70   981.15   0.848   1.000  29.347       41.51

15   80.000   84.980     1.55   894.50   0.833   1.000  29.849       34.52

16   85.000   89.980     1.45   819.34   0.818   1.000  30.321       29.48

17   90.000   94.980     1.39   754.11   0.803   1.000  30.766       25.87

18   95.000   99.980     1.36   697.47   0.788   1.000  31.189       23.34

19  100.000  104.980     1.36   648.26   0.774   1.000  31.590       21.62

20  105.000  109.980     1.40   605.50   0.760   1.000  31.971       20.54

21  110.000  114.980     1.46   568.34   0.747   1.000  32.336       19.99

22  115.000  119.980     1.54   536.07   0.735   1.000  32.685       19.89

23  120.000  124.980     1.67   508.08   0.723   1.000  33.019       20.22

24  125.000  129.980     1.82   483.86   0.713   1.000  33.340       20.96

25  130.000  134.980     2.02   462.97   0.703   1.000  33.648       22.16

26  135.000  139.980     2.28   445.05   0.694   1.000  33.945       23.90

27  140.000  144.980     2.61   429.80   0.685   1.000  34.231       26.34

28  145.000  149.980     3.05   416.94   0.678   1.000  34.508       29.74

29  150.000  154.980     3.64   406.29   0.672   1.000  34.775       34.58

30  155.000  160.000     4.50   397.63   0.667   1.000  35.034       41.83

Block  #      Angle - Limits    Average     Total    Step Time

#   pts     Low       High   Int./1000    Time     sec/step

1   250    10.000    14.980   4884.23    362.04      1.45

2   250    15.000    19.980   2563.81    384.24      1.55

3   250    20.000    24.980   1493.85    417.72      1.65

4   250    25.000    29.980    930.14    466.34      1.85

5   250    30.000    34.980    607.34    534.82      2.15

6   250    35.000    39.980    411.06    629.05      2.50

7   250    40.000    44.980    286.19    756.26      3.05

8   250    45.000    49.980    204.00    924.97      3.70

9   250    50.000    54.980    148.50   1144.49      4.60

10   250    55.000    59.980    110.30   1424.02      5.70

11   250    60.000    64.980     83.59   1771.14      7.10

12   250    65.000    69.980     64.70   2189.77      8.75

13   250    70.000    74.980     51.21   2677.71     10.70

14   250    75.000    79.980     41.51   3224.28     12.90

15   250    80.000    84.980     34.52   3809.10     15.25

16   250    85.000    89.980     29.48   4402.68     17.60

17   250    90.000    94.980     25.87   4969.51     19.90

18   250    95.000    99.980     23.34   5473.17     21.90

19   250   100.000   104.980     21.62   5881.78     23.55

20   250   105.000   109.980     20.54   6172.39     24.70

21   250   110.000   114.980     19.99   6332.92     25.35

22   250   115.000   119.980     19.89   6361.64     25.45

23   250   120.000   124.980     20.22   6264.97     25.05

24   250   125.000   129.980     20.96   6054.63     24.20

25   250   130.000   134.980     22.16   5745.04     23.00

26   250   135.000   139.980     23.90   5351.25     21.40

27   250   140.000   144.980     26.34   4887.76     19.55

28   250   145.000   149.980     29.74   4367.80     17.45

29   250   150.000   154.980     34.58   3803.11     15.20

30   251   155.000   160.000     41.83   3215.39     12.80

Total time = 100012.80

## Data analysis

It is important that the data analysis software used to process VCT data be modified to handle the new format. with the step counting time and the step intensity being read in at each point. The step counting time is used to modify the calculated step intensities, not to simply scale the observed intensities. Since the weighting scheme is usually based on the accumulated count at each point, the correct weighting will be maintained.

## References

Madsen I.C. and Hill R.J. (1991) “Rietveld Analysis Using Para-Focusing and Debye Scherrer Geometry Data Collected with a Bragg-Bentano Diffractometer”, Zeitschrift fur Kristallographie, 196, 73-92

Madsen I.C. and Hill R.J. (1992), “Variable Step-Counting Times For Rietveld Analysis, or, Getting The Most Out of Your Experiment Time”, Advances in X-ray Analysis, 35, 39-47

Madsen I.C. and Hill R.J. (1994), “Collection and Analysis of Powder Diffraction Dtaa with Near Constant Counting Statistics”, J.Appl. Cryst, 27, 385-392