## Slide 1
Quantitative Analysis (QXRD)
David Hay
CSIRO Manufacturing Science and Technology, Locked Bag 33,
Clayton South MDC, Clayton 3169
Recommended Text:
“Quantitative X-Ray Diffractometry”
Lev S.Zevin, Giora Kimmel
ed Inez Mureinik
Springer-Verlag, New York, 1995
Cost: approx. $99 A
Notes: 1

## Slide 2
OUTLINE
Nomenclature
Introduction/Background
Basic Concepts
Techniques
  Absorption-Diffraction
  Internal Standard Analysis (RIR)
  Doping or Spiking method
  Dilution method
  Complete n-phase analysis (external standard method)
  Whole pattern profile matching
Notes of interest
  Rapid method using PSD (Cressey)
  Matrix method (Knudsen)
Notes: 2

## Slide 3
NOMENCLATURE (consistent with Klug and Alexander)
i,j,k, … (lowercase subscript) lines in diffraction pattern
I,J,K, …(uppercase subscript) components of a mixture
S subscript referring to component used as reference standard
ρJ Density of component J
fJ Volume fraction of component J in a mixture
xJ Weight fraction of component J in a mixture
μJ Linear absorption coefficient of component J for x-ray wavelength λ
μJ* = μJ / ρJ Mass absorption coefficient of component J for x- ray wavelength λ
Notes: 3

## Slide 4
μ Linear absorption coefficient of sample with a number, N of components
μ* Mass absorption coefficient of sample with a number, N of components
N
μ* = ∑ xJ μJ*
J=1
IiJ Intensity of ith line of component J in mixture
(IiJ)0 Intensity of ith line of pure component J
I / Icor Ratio of intensity of strongest peak of each component to a common standard, synthetic corundum, in a 1:1 mixture of sample and standard
RIRJ,S Reference Intensity Ratio of phase J with respect to phase S
Notes: 4

## Slide 5
INTRODUCTION AND BACKGROUND
Powder XRD ideal for crystalline mixture analysis:
  Each component produces characteristic pattern superimposed on those of the other components – unscrambling these superimposed patterns leads to identification of the components
  Intensity of lines in each components pattern is proportional to the concentration of that component, disregarding the effects of absorption.
Quantitative analyses first undertaken in the 1920’s and 30’s (eg Clarke and Reynolds developed scheme for mine dust analysis using an internal standard procedure).
Notes: 5

## Slide 6
The first Geiger-counter diffractometers were used in 1945 and these allowed data of sufficient precision to be recorded for QA.
Absorption: When a mixture contains both a weak and a strong absorber, lines of the weak absorbing compound appear weaker (and those of the strong absorbing compound stronger) than calculated from a linear relationship between pattern intensity and composition.
In 1948 Alexander and Klug gave the theoretical mathematical background for the effects of absorption on diffracted intensities from a flat powdered cake. Many developed methods have used their basic equations.
Notes: 6

## Slide 7
Issues
Note: From IUCr CPD Round Robin on Quantitative Phase Analysis
(http://www.dl.ac.uk/SRS/XRD/IUCR/QARR/intro.htm)
“The round robin will address the following
analytical issues:”
Type of analysis
diffraction (X-ray /Neutron) vs non- diffraction
internal std vs external std vs spiking etc
standardless methods
Notes: 7

## Slide 8
Sample features
representivity & homogeneity
particle & crystallite size
statistics & microabsorption
crystallinity & surface roughness
preferred orientation, microabsorption & extinction
Data collection
type of instrument / geometry (eg Bragg-Brentano, Seemin- Bohlin, Guinier, reflection or transmission, Debye Scherrer)
sample preparation
data range and wavelength
Notes: 8

## Slide 9
Data analysis
integrated intensities vs full-profile
Rietveld vs data-base of observed patterns
use of constraints and corrections
software systems & methods
complexity of the pattern - peak overlap
Notes: 9

## Slide 10
BASIC CONCEPTS
Integrated intensity proportional to analyte phase concentration
Intensity functionally dependent on absorbing power of mixture for X-rays
Dependence is generally non-linear and unknown
Absorption
IiJ = K ‘iJ fJ / μ = K ‘iJ xJ / ρJ μ*
Density is constant for given phase so can be included in K, so simplified to:
n
IiJ = KiJ xJ / μ* = KiJ xJ / Σ xK μK* … (1)
K=1
Notes: 10

## Slide 11
IiJ = KiJ xJ / μ* … (1)
Object is to determine xJ … achieve this by:
Measuring intensity of IiJ
Determining mass absorption coefficient of sample μ*
Determining calibration constant KiJ
Notes: 11

## Slide 12
TECHNIQUES
1. Absolute techniques where the intensity of the peak in the pure analyte (IiJ)0 is involved in the equations to solve for xJ Note these techniques depend on experimental conditions (instrument type, settings etc) so it is difficult to compare results from different laboratories.
2. Relative or Ratio techniques where ratios, not absolute intensities, are used. (eg internal and external standard techniques). Independent of experimental conditions.
Notes: 12

## Slide 13
TECHNIQUES
6 Basic Techniques to be described
Absorption-Diffraction … known mass absorption coefficient for mixture
Internal Standard … sample doped with reference
Doping (Spiking) … sample spiked with the analyte phase
Dilution … sample diluted with inert diluent
Complete n-phase Analysis
Whole Pattern Profile Matching
Notes: 13

## Slide 14
TECHNIQUE #1 Absorption-Diffraction
Known Mass Absorption Coefficient … (absolute class)
μ* previously determined by:
  calculation from known composition, summation of atomic absorption coefficients
  direct absorption measurement
  measurement of Compton scattering
  case with constant μ* (includes phase transitions where chemical composition remains unchanged during the transition)
Notes: 14

## Slide 15
General case:
multiple phases, variable absorption coefficient
  requires measurement of (IiJ)0 for the pure phases in the mixture
  applying … (1) to the pure phase (ie xJ = 1) (IiJ)0 is defined:
(IiJ)0 = KiJ / μ* J
  By combining this with … (1) Leroux et al (1953) obtained expression for phase abundance:
xJ = IiJ / (IiJ)0 • μ* / μ* J …(2)
Notes: 15

## Slide 16
Suggested measurement of μ* by direct transmission through the specimen … hence called diffraction-absorption technique
Leroux’s equation can be re-written:
log xJ - log [IiJ / (IiJ)0 ] = log [μ* / μ* J]
Notes: 16

## Slide 17
Plot of log xJ -log [IiJ / (IiJ)0 ] vs log [μ* / μ* J] at constant x should be straight line of slope -1
In practice large deviations because of microabsorption# so … (2) becomes:
xJ = IiJ / (IiJ)0 • [μ* / μ* J] α
where α is the slope of the regression line of log [μ* / μ* J] vs log xJ -log [IiJ / (IiJ)0 ]
# NB Microabsorption occurs when two substances of different mass absorption coefficients are mixed.
Notes: 17

## Slide 18
α is not constant but depends on grain size of analyte and the nature of the matrix of other phases.
Note that microabsorption effects can be severe in this method due to the possibility of the pure phase J and the multiphase sample having different absorption properties and crystallite size distributions.
Precision: Approximately 5% relative error
Detection Limit: Measured in tenths of percent
(NB precision and LOD as quoted by Zevin are probably optimistic)
Notes: 18

## Slide 19
TECHNIQUE #2 Internal Standard Analysis
(relative class)
Sample doped with known amount of reference material (internal standard)
  The weight fraction of the unknown phase J in the doped sample x’J is given by:
x’J = xJ (1-xS)
where xJ is weight fraction of phase J in the undoped sample, xS is the weight fraction of the internal standard.
Notes: 19

## Slide 20
  Phase J and the internal standard must have at least one resolved diffraction peak then:
IiJ = KiJ xJ (1- xS ) / μd* (for phase J)
IhS = KS xS / μd* (for internal standard)
Note: μd* is mass absorption coefficient of doped sample
  Therefore, the ratio of intensities is independent of absorption:
IiJ / IhS = KiJ xJ(1-xS) / KhS xS …(3)
Notes: 20

## Slide 21
  From (3) weight fraction xJ is a linear function of intensity ratio:
xJ = [xS / (KiJ / KhS ) (1-xS) ] • [IiJ / IhS ]
= constant • [IiJ / IhS ]
  Internal standard method can be used to analyze any phase in a mixture without analyzing remaining phases. If all phases determined, then mass balance will hold.
Precision: Approximately 1% relative error
Detection Limit: Measured in tenths of percent
Notes: 21

## Slide 22
Reference Intensity Ratio
Easiest way to bring diffraction peak intensities to common scale is to scale them to a particular peak in a common reference phase.
deWolff and Visser (1964) suggested the (113) peak in corundum (100% line).
  Hence, Reference Intensity Ratio (RIR) is ratio of strongest peak of phase J to that of the corundum (113) in a 1:1 (w/w) mixture.
RIRJ,c = Ii,J / Ih,c
  Extension to general case RIR of phase α wrt phase β
RIRα,β = Ii,α / Ih,β
Notes: 22

## Slide 23
Reference Intensity Ratio
  Usually quoted as ratio of direct peak heights but integrated intensities also used
  Any number of constituents may be quantified
  Mixture may contain amorphous components
  Mass absorption coefficient of mixture need not be known in advance
Notes: 23

## Slide 24
TECHNIQUE #3 Doping or Spiking Method
(absolute class)
Sample to be analysed is spiked with known amount of the analyte phase J
  Weight fraction of phase J in spiked sample
M = mass sample MJD = mass dopant phase J xJM = mass of phase J in undoped sample xJD = weight fraction of dopant = MJD / (M + MJD)
(xJ)D = (xJM + MJD) / (M + MJD) = xJ (1 - xJD) + xJD
Notes: 24

## Slide 25
  Mass absorption coefficient of doped sample
Mass absorption coefficient of doped sample, μD* , is the sum of the ratios of the doped fraction xJD , with absorption coefficient μJ*, and the remaining fraction (1- xJD) with absorption coefficient μ*
μD* = μ* (1 - xJD ) + μJ* xJD
Now substitute these values for (xJ)D and μD* into …(1)
(IiJ)D = KiJ [xJ (1-xJD) + xJD] / [μ* (1- xJD) + μJ xJD ]
= KiJ [xJ (1-xJD) + xJD] / [μ* {1+ xJD(μJ* - μ* ) / μ* }]
…(4)
Notes: 25

## Slide 26
In some cases the dopant will not change the overall μ* significantly. i.e. μ* ≈ μD*
Then
(IiJ)D = KiJ [xJ (1-xJD) + xJD] / [μ* {1+ xJD(μJ* - μ* ) / μ* }]
can be simplified to
(IiJ)D = KiJ [xJ (1-xJD) + xJD] / μ*
= (KiJ / μ* ) xJ + (KiJ / μ* ) (1 - xJ) xJD
Notes: 26

## Slide 27
Then for a series of dopant concentrations, a plot of (IiJ)D vs dopant concentration, xJD , is linear
(IiJ)D
xJD
0
x0
IiJ
At x0, (IiJ)D = 0 and xJ = x0 (1- x0)
Notes: 27

## Slide 28
However in most cases μ* is not constant
In this case, can use …(1) for the undoped sample
and …(4) for the doped sample
ie two equations with three unknowns, viz.
xJ , KiJ and μ*
We can introduce a third equation from the pure dopant phase viz.
(IiJ)0 = KiJ / μ J *
and so have three equations in three unknowns.
Notes: 28

## Slide 29
TECHNIQUE #4 Dilution Method
(absolute class)
Sample diluted with unreactive material of known μ*
  Diluent may be either crystalline or amorphous but should not contain phases to be analysed
  Instead of measuring sample absorption, measurements of diluted and undiluted samples are made
  Diffraction from diluent must not interfere with analytical diffraction peak
  Diluent should not fluoresce, should be chemically inert, exist as fine powder and have mass absorption coefficient close to that of the sample
Notes: 29

## Slide 30
Weight fraction diluent (D)
xD = MassD / (M + MassD) where M = Mass original sample
Weight fraction phase J in diluted sample
xJ M / (M + MassD ) = xJ (1 - xD ) see Zevin, p130 for derivation
Mass absorption coefficient of diluted sample
(μ )D* = μD* xD + (1 - xD) μ * rule of additivity
We can now use ...(1) to make three equations in three unknowns (μ * , Ki,J and xJ ):
(i) Ii,J = Ki,J xJ / μ *
(ii) (Ii,J )0 = Ki,J / μ J * ie when xJ = 1, ie pure phase J
(iii) (Ii,J )D = Ki,J xJ(1-xD) / [μ D*xD + (1 - xD) μ * ]
Notes: 30

## Slide 31
Solving for xJ gives:
xJ = [ xD μD* / (1-xD) μJ* ] • [ ( (Ii,J)D / (Ii,J)0 ) / (1-((Ii,J)D / Ii,J )]
Precision and accuracy
  Optimal dilution gives ( Ii,J )D / Ii,J approximately 0.33
  This gives precision of approximately 2%-3% with relative error of analysis of about 10%.
Notes: 31

## Slide 32
TECHNIQUE #5 Complete n-phase analysis
(relative class)
For analysis of all phases, an additional mass balance equation is available:
∑ xJ = 1
J
Two possibilities:
1 Determine weight fractions of each phase by technique #1, #2, #3, #4 etc. Then use mass balance equation to determine accuracy of analysis. ie deviation from a sum of 1 indicates either systematic error or misses phase(s).
Notes: 32

## Slide 33
2 The mass balance equation can be used to exclude one unknown (eg μ * ) from calculation of weight fractions.
e.g. Equation … (1) can be rearranged as:
xJ = μ* I iJ / KiJ
If both sides are summed over all the phases in the sample then:
n
1 = μ* ( Σ I iJ / KiJ)
J=1
or:
n
μ* = ( Σ I iJ / KiJ) -1
J=1
Notes: 33

## Slide 34
By combining our rearranged equation … (1) with that for μ* we can obtain an expression for any phase, let’s say R, free of the absorption coefficient:
n
xR = [ IhR / KhR ] / [ Σ (IiJ / KiJ) ]
J=1
n
xR = 1 / Σ ( βRJ / SRJ) … (4)
J=1
where: SRJ is the ratio of intensities of peaks belonging to phases J and R respectively ie SRJ = IiJ / IhR
and βRJ = KhR / KiJ
Notes: 34

## Slide 35
  The feature of this method that makes it highly attractive is that equation … (4) contains only ratios and no absolute values of intensities and calibration constants. Thus the method can be considered independent of measurement conditions.
  Determination of βRJ ratios by matrix methods (see Zevin, p136 ff).
  Also called external standard method (Hubbard C.R., Evans E.H, Smith, D.K., J.Appl.Cryst., 9, 169-74, 1976) since the matrix analysis utilises a third phase K not found or doped in the sample in acquiring the βRJ ratios.
Notes: 35

## Slide 36
TECHNIQUE #6 Whole Pattern Profile Matching
  Most useful where overlapping peaks are a problem .. Phase to be analysed may not have at least one clearly resolved peak.
Profile fitting Groups of overlapping peaks are separated into individual peaks and integrated intensities extracted. Knowledge of peak profiles, positions required but not crystal structure or composition.
Notes: 36

## Slide 37
Structure Refinement Either using the conventional methods developed for single crystal structural analysis (extraction of integrated intensities, refinement of derived structure factors against a model) or by continuous pattern Rietveld methods.
  Both structure refinement methods require knowledge of the crystal structure
Notes: 37

## Slide 38
  Numerous references on QXRD using Rietveld methods, with particular reference to work of Hill, Howard, Madsen (eg Hill, R.J., Howard, C.J., J.Appl.Crystallogr., 20, 1987, 467-74.)
Ik = S Mk LPk ⎪F2⎪
S is the Rietveld scale factor
n
xJ = [ SJ ZMVJ ] / [ Σ Si ZMVi ]
i=1
Z = number of formula units per unit cell M = atomic mass of the formula unit V = volume of one unit cell
Notes: 38

## Slide 39
NOTES OF INTEREST
1. Absorption and Microabsorption
G. Cressey and M. Batchelder in IUCr Newsletter No. 20
http://www.iucr.org/iudr-top/comm/Newsletters/
Uses PSD … fits the mixed assemblage pattern with proportionally reduced 100% phase patterns of identified components.
Overcomes problems with preferred orientation, absorption and micro-absorption (from surface roughness).
Notes: 39

## Slide 40
2. Quantitative Analysis with Qualitative Control of Calibration Standards
T. Knudsen, X-Ray Spectrometry, 10 (2), 54-6, 1981
“Samples containing identical or very similar phases and differing in their contents of these can be used for calibration of a quantitative X-ray diffraction analysis without knowing their actual mineralogical composition. The method requires that at least one major peak for each phase is well isolated and that the number of calibration samples is equal to, or exceeds, the number of phases”
Notes: 40

## Slide 41
For a system of m phases and n samples (where n ≥ m), and sum of weight % of all phases is 100%, a set of equations is derived of the form:
100 = K1I11 + K2I12 + … + Km I1m
100 = K1I21 + K2I22 + … + Km I2m
•
•
•
100 = K1In1 + K2In2 + … + Km Inm
ie this is a matrix of the form L = IK
with L and K being column matrices with the elements 100 and the KJ ‘s respectively, and I a rectangular matrix containing the IiJ ‘s.
Notes: 41

## Slide 42
100
100
•
•
•
100
I11 I12 • • • I1m
I21 I22 • • • I2m
•
•
•
In1 In2 • • • Inm
K1
K2
•
•
•
Km
L = I K
Notes: 42

## Slide 43
A least squares solution for the calibration constants KiJ is given by:
K = (I T I ) -1 I T L
with the superscripts T and -1 representing the transpose and inverse matrix.
This operation can be conveniently carried out using a spreadsheet (eg Microsoft EXCEL).
Notes: 43

## Slide 44
BIBLIOGRAPHY
L.E.Alexander and H.P.Klug, Anal. Chem., 20,886,1948
H.P.Klug and L.E.Alexander “X-Ray Diffraction Procedures for Polycrystalline and Amorphous Materials”, 2nd Edition, John Wiley NY, 1974, ISBN 0-471-49369-4
L.S.Zevin and G.Kimmel, “Quantitative X-Ray Diffractometry”, Springer NY, 1995, ISBN-0-387-94541-5
R.Jenkins and J.L.deVries, “An Introduction to X-Ray Powder Diffractometry”, N.V.Philips, Eindhoven
G.Cressey and M.Batchelder, “Dealing with Absorption and Microabsorption in Quantitative Phase Analysis”, IUCr Newsletter No 20., http://www.iucr.org/iucr-top/comm/cpd/Newsletters/
T.Knudsen, “Quantitative X-Ray Diffraction Analysis with Qualitative Control of Calibration Samples”, X-Ray Spectrometry, 10(2), 54-6, 1981.
R.J.Hill & C.J.Howard, J.Appl.Crystallogr., 20, 467-74, 1987
G.L.Clark & D.H.Reynolds, Ind.Eng.Chem., Anal.Ed., 8,36,1936
J.Leroux, D.Lennox, D.Kay, Anal.Chem.,25, 740-3, 1953
P.M.deWolff & J.W.Visser, Technisch Pysische Dienst. Rept No 641, Delft, 1964
Notes: 44